Bruker 3-Letter codes

Messung starten
gilt für http://localhost/OpusCommand.htm?MeasureSample und
http://localhost/OpusCommand.htm?MeasureReference

EXP='frank.xpm', = Experiment mit allen Parametern abgespeichert
XPP='C:\OPUS_7.0.129\XPM', = Pfad zum Experiment
ADT=0, = Additional Data Treatment (do not change)
APF='B3', = Apodization Function Blackaman Harris 3 Term
  BX = Boxcar
  TR = Triangular
  HG = Happ-Genzel
  B3 = Blackman-Harris 3-Term
  B4 = Blackman-Harris 4-Term
  NBW = Norton-Beer, Weak
  NBM = Norton-Beer, Medium
  NBS = Norton-Beer, Strong

APT='Open', = Aperture Open
 Open = Blende offen
  noch zwei, drei String Einträge die ich im Moment nicht nachschlagen kann, eventuell auch beim OPF.Parameter, Info kommmt noch

AQM='DD', = Akquisition Mode Double Sided, backward + forward
  SN = Single Sided
  DN = Double Sided
  SF = Single Sided Fast Return
  SD = Single Sided, Forward-Backward
  DD = Double Sided, Forward-Backward
  DF = Double Sided, Fast Return

BMS='CaF2', = Beam Splitter CaF2 (bleibt gleich)
BSF='C:\OPUS_7.0.129\', = Background Spectrum File ' (not used)
CHN='Module port', = Measurement Channel „Module Port“
CNM='Administrator', User Administrator (must also exist in Bruker software)
DEL=0, = Delay Before Measurement
DIG=0, = Digital Filter
DLY=0, = Stabilization Delay
DMX='1', = Multiplexed Data
DTC='TE-MCT-PV [Internal]', = Detector settings (keep always)
GSG='ON', = Gain Switch Gain
HFQ=1800.000000, = End Frequency Limit for File in Hz (integer)
HFW=8000.000000, = Wanted High Frequency Limit in Hz (integer)
IRS=700, = Iris Aperture (micron)
HPF=-1, High Pass Filter
LFQ=4000.000000, = Start Frequency Limit for File in Hz (integer)
LFW=1800.000000, = Wanted Low Frequency Limit in Hz (integer)
LPF='-1', Low Pass Filter
MIN=1.000000, = Sample Meas. Duration in Min.
MIR=1.000000, = Background Meas. Duration in Min.
NAM='<snm>', = Filename <snm> increments name SNM (=Sample Name) automatically
NLI=1, = Non Linearity Correction
NSS=32, = Sample Scans , ganze Zahl
NSR=32, = Background Scans , ganze Zahl
OPF='Open', = Optical Filter Setting
PGN='1', = Preamplifier Gain
PGR='1', = Preamplifier Gain Background
PHR=32.000000, = Phase Resolution , , ganze Zahl
PHZ='PW', = Phase Correction Mode
  PW = Power Spectrum

PLF='AB', = Result Spectrum
  TR = Transmittance
  AB = Absorbance

PTH='C:\OPUS_7.0.129\MEAS3\', = Measurement Path
RCH='Module port', = Reference Channel „Module Port“
RES=1.000000, = Resolution in cm-1 (number between 1 and 10)
RGN='-1', = Signal Gain Background -1 (Automatic)
SFM='MATRIX-MF', = Sample Form (always stays at this value)
SGN='-1', = Signal Gain -1 (Automatic)
SNM='Test24', = Sample Name
SON='Off', = External Synchronisation
VEL='40.0', = Scanner Velocity
ZFF='4' = Zero Filling for FFT, values my be 0, 2, 4, 8 (power of 2)

Messung speichern (Bem.: auch die HTTP-Anwort der Messung auswerten siehe ReadMe_html.txt)
http://localhost/OpusCommand.htm?SaveAs ...

DAP='C:\OPUS_7.0.129\MEAS3', Pfad zu den gespeicherten files (Data File Path)
OEX='0', (OEX overwrite existing file = 0 no!) bleibt
SAN='Test24.1.dpt', Name des files ('Save As' File Name)
COF=64, Copy Flags
DPA=6, Anzahl der Stellen hinter dem Komma für Wellenzahl (Decimal Places, Abscissa)
DPO=6, Anzahl der Stellen hinter dem Komma für Extinktion (Decimal Places, Ordinate)
SEP=',', Separator für Spalten in File
YON='0', Y-Values only
ADP='1', All data points
X64='0'

Konfiguration
Einteilung in 4 Bereiche

OPTIK, Hardware
AQM, BMS, APT, OPF, RES, LFQ, LFW

FFT, Calculation
APF, PHR, PHZ, PLF

EXPERIMENT
CNM, EXP, XPP, PTH, NAM
NSS, NSR, HFQ, HFW

SAVE
DAP, SAN, DPA, DPO