Measurement
profile

### preparation
- prerequisite: instrument is on and stabilized
- activate nafion dryer (always do before using chemical (perchlorate) dryer (must run at least ½ hour)!!!
- configure the measurement (measurement type one/many/continuously, calibrate with HITRAN or standard gases, parameters to accumulate spectra and calculate FFT, and parameters to calculate mixing ratios with OCTAVE)
- prepare the list of actions

### execution
actions are one of the following
1. purge with N2 gas (or H20 and CO2 free gas) for reference measurement
2. do reference measurement
3. if necessary do calibrations on one/some gas components (one aftet the other)
4. inlet of sample through nafion and perchlorate filters to reduce H2O
5. do measurement (one/many/contiuous) on sample(s)

- repeat reference and calibration measurements as often as desired (specified in list of actions)
- pressurize cell for higher sensitivity if desired (then do all measurements/references /calibrations with same pressure)

### output is mixing ratio ppm or µMol/Mol of selected gas components
- display of actual values online
- display of time course and Allen plots

### performed in background
- calculation of mixing ratios out of spectra
- calibrated by HITRAN data base or user supplied reference/standard gas